## Sheet1
400 CORE
자석패키지 | #N/A | 자석사이즈 | #N/A
스탠드패키지 | #N/A
최소 천장 높이+800 | #N/A
최소 문 폭(magnet+50) | #N/A
자석 중량 | #N/A | 전체 중량 | #N/A
스탠드 중량 | #N/A
설치질소 | #N/A | 질소볼륨 | #N/A
설치헬륨 | #N/A | 헬륨볼륨 | #N/A
질소리필주기 | #N/A | 질소소모량 | #N/A
헬륨리필주기 | #N/A | 헬륨소모량 | #N/A

## Magnet
magnet | 자석패키지 | 스탠드패키지 | 자석 중량 | 스탠드 중량 | 전체 중량 | 자석사이즈 | 최소 천장 높이+80 | 최소 문 폭(magnet+5) | 설치질소 | 설치헬륨 | 질소볼륨 | 헬륨볼륨 | 질소리필주기 | 헬륨리필주기 | 질소소모량 | 헬륨소모량
400core | 1250*1040*1580 | 790*760*1200 | 508 | 154 | 590 | 1285*1154 (지름=1430) | 2178+800 | 90 | 450 | 300 | 102 | 150 | 16days | 365days | 230ml/h | 13ml/h
400evo | 1250*1040*1581 | 790*760*1201 | 508 | 154 | 590 | 1285*1154 (지름=1430) | 2178+800 | 90 | 450 | 300 | 102 | 150 | 16days | 365days | 230ml/h | 13ml/h
500evo | 1250*1040*1582 | 790*760*1202 | 588 | 154 | 660 | 1285*1154 (지름=1430) | 2178+800 | 90 | 500 | 350 | 102 | 134 | 16days | 365days | 230ml/h | 13ml/h
600evo | 1250*1040*1583 | 790*760*1203 | 612 | 168 | 700 | 1285*1154 (지름=1430) | 2308+800 | 90 | 500 | 450 | 102 | 129 | 16days | 365days | 230ml/h | 13ml/h
700evo | 1250*1040*1584 | 790*760*1204 | 882 | 174 | 950 | 1285*1154 (지름=1430) | 2508+800 | 90 | 600 | 500 | 128 | 138 | 16days | 365days | 260ml/h | 17ml/h

## Console
콘솔 | 장비 사이즈 | 장비 무게 | 필요전기 | 필요가스
OneBay | 690*830*1300 | 230 | 3.8kW | 60Nl/min
TwoBay | 1310*830*130 | 340 |  | 60Nl/min
Nanobay | 450*880*710 | 130 | 3.8kW | 60Nl/min

## Auto
AutoSampler
| 장비 사이즈 | 장비 무게 | 필요전기 | 필요가스
Sample Case 24 |  |  |  | 45Nl/min
Sample Case Plus |  |  |  | 45Nl/min
Sample Case Heated & Cooled |  |  |  | 45Nl/min
Sample Jet |  |  |  | 45Nl/min

## CPP&CRP
| 장비 사이즈 | 장비 무게 | 전력소모 | 추천 | 필요전기 | 차단기용량 | 평상시 | 가스소모 | 쿨링라인 | 플로우레이트 | 쿨링워터온도
Prodigy | 600*400*460 | 68 | 0.6kW
LN2dewar | 500*620*1350 |  |  |  | 220V |  |  | 100L
CU | 740*740*980 | 250 | 1kW,0.3kW |  | 220V |  |  | 1L/min
Outdoor | 930*390*950 | 142 | 9.0kW, 8.0kW | 12kW | 460 |  | 13A |  | 3/4"
indoor | 270*600*450 | 46 |  |  |  | 20A~30A
Water Cooled | 450*490*590 |  | 8.5kW, 7.5kW | 12kW | 460 | 17A~30A |  |  |  | 6~10L/Min | 20도 미만
Helium Gas Transfer Lines |  |  |  | Helium Gas Lines
For Indoor Helium Compressors (3 Options)
• 6 m / 10 m / 20 m |  |  |  | • 10 m ~60 m
For Outdoor Helium Compressors
• 3 m indoor line / 10 m outdoor line
• 3 m indoor line / 20 m outdoor line
• 3 m indoor line / 30 m outdoor line
• 6 m indoor line / 10 m outdoor line
• 6 m indoor line / 20 m outdoor line
• 6 m indoor line / 30 m outdoor line
• 10 m indoor line / 10 m outdoor line
• 10 m indoor line / 20 m outdoor line
• 10 m indoor line / 30 m outdoor line
• 20 m indoor line / 10 m outdoor line
• 20 m indoor line / 20 m outdoor line
• 30 m indoor line / 10 m outdoor line

## ETC
| 장비 사이즈 | 장비 무게 | 필요전기 | 필요가스
MAS |  |  |  | 250Nl/min
BCU1 | 480*350*430 | 30
BCU2 | 580*420*570 | 70